## Problem
Page 1: Coagulation PT Reagent New Lot Validation
Date | 20/4/07 | Analyzer | TED | New Lot# | 12345 | Old Lot# | 12344 | Yes
Tech | EPM | Serial Number | abcde | Expiration | 2008-09-30 | Expiration | 2007-05-31 | No
|  |  |  | ISI | 1.1 | ISI | 1.05
Current Geometric Mean | 11.3 | Current Reference Range: |  | Low: | 10.1 | High: | 12.7
Results from testing |  |  |  |  | Manufacturer's Expected Percent Difference |  |  | 2.5
Run | Old Lot |  | New Lot
| PT Result | INR | PT Result* | INR | * Results outside current reference range shown in bold |  |  | Difference | Percent
1 | 10.9 | 0.96287 | 12.2 | 1.0854 |  |  |  | -1.3 | -11.9266
2 | 11 | 0.97214 | 11.2 | 0.98794 |  |  |  | -0.2 | -1.8182
3 | 11.3 | 1 | 11.5 | 1.0171 |  |  |  | -0.2 | -1.7699
4 | 11.2 | 0.99071 | 10.8 | 0.94919 |  |  |  | 0.4 | 3.5714
5 | 11.9 | 1.0558 | 10.9 | 0.95887 |  |  |  | 1 | 8.4034
6 | 11.4 | 1.0093 | 10.6 | 0.92988 |  |  |  | 0.8 | 7.0175
7 | 12.1 | 1.0745 | 11.7 | 1.0366 |  |  |  | 0.4 | 3.3058
8 | 11.7 | 1.0372 | 12.7 | 1.1344 |  |  |  | -1 | -8.547
9 | 12.4 | 1.1025 | 10.1 | 0.88174 |  |  |  | 2.3 | 18.5484
10 | 11.3 | 1 | 11.3 | 0.99764 |  |  |  | 0 | 0
11 | 12.7 | 1.1305 | 12 | 1.0658 |  |  |  | 0.7 | 5.5118
12 | 11.9 | 1.0558 | 11.3 | 0.99764 |  |  |  | 0.6 | 5.042
13 | 10.8 | 0.95359 | 10.9 | 0.95887 |  |  |  | -0.1 | -0.92593
14 | 11.4 | 1.0093 | 11.1 | 0.97824 |  |  |  | 0.3 | 2.6316
15 | 11 | 0.97214 | 11 | 0.96855 |  |  |  | 0 | 0
16 | 11.8 | 1.0465 | 11.6 | 1.0268 |  |  |  | 0.2 | 1.6949
17 | 11.7 | 1.0372 | 12 | 1.0658 |  |  |  | -0.3 | -2.5641
18 | 11.4 | 1.0093 | 11 | 0.96855 |  |  |  | 0.4 | 3.5088
19 | 12.7 | 1.1305 | 11.9 | 1.0561 |  |  |  | 0.8 | 6.2992
20 | 12.6 | 1.1211 | 11 | 0.96855 |  |  |  | 1.6 | 12.6984
New Geometric Mean |  |  | 11.3243
Current Refrence Range Validated? |  |  |  | Yes
If no, complete page 2
Comments: Unacceptable. Repeat
Slope: | 0.10879
R: | 0.11195
Intercept: | 10.4263
Supervisor Review EL 1/5/07
Page 2: Updated Reference Ranges
|  | Date |  | Analyzer |  | New Lot#
|  | Tech |  | Serial Number |  | Expiration
| Prothromin Time |  |  |  | INR
| Old Lower Limit |  | Old Upper Limit |  | Old Lower Limit |  | Old Upper Limit
| New Lower Limit |  | New Upper Limit |  | New Lower Limit |  | New Upper Limit
New Adverse Event Limits
|  | PT | INR |  | Results greater than the values listed are Adverse Events as graded
| Grade 1 | 0 | 0
| Grade 2 | 0 | 0
| Grade 3 | 0 | 0
| Grade 4 | 0 | 0
Supervisor Review

## Good
Page 1: Coagulation PT Reagent New Lot Validation
Date | 20/4/07 | Analyzer | TED | New Lot# | 12345 | Old Lot# | 12344 | Yes
Tech | EPM | Serial Number | abcde | Expiration | 2008-09-30 | Expiration | 2007-05-31 | No
|  |  |  | ISI | 1.1 | ISI | 1.05
Current Geometric Mean | 11.3 | Current Reference Range: |  | Low: | 10.1 | High: | 12.7
Results from testing |  |  |  |  | Manufacturer's Expected Percent Difference |  |  | 2.5
Run | Old Lot |  | New Lot
| PT Result | INR | PT Result* | INR | * Results outside current reference range shown in bold |  |  | Difference | Percent
1 | 10.9 | 0.96287 | 10.8 | 0.92203 |  |  |  | 0.1 | 0.91743
2 | 11 | 0.97214 | 11.1 | 0.95024 |  |  |  | -0.1 | -0.90909
3 | 11.3 | 1 | 11.5 | 0.98798 |  |  |  | -0.2 | -1.7699
4 | 11.2 | 0.99071 | 11 | 0.94083 |  |  |  | 0.2 | 1.7857
5 | 11.9 | 1.0558 | 12.1 | 1.0448 |  |  |  | -0.2 | -1.6807
6 | 11.4 | 1.0093 | 11.5 | 0.98798 |  |  |  | -0.1 | -0.87719
7 | 12.1 | 1.0745 | 12 | 1.0353 |  |  |  | 0.1 | 0.82645
8 | 11.7 | 1.0372 | 11.7 | 1.0069 |  |  |  | 0 | 0
9 | 12.4 | 1.1025 | 12.3 | 1.0638 |  |  |  | 0.1 | 0.80645
10 | 11.3 | 1 | 11.3 | 0.96909 |  |  |  | 0 | 0
11 | 12.7 | 1.1305 | 12.5 | 1.0829 |  |  |  | 0.2 | 1.5748
12 | 11.9 | 1.0558 | 11.8 | 1.0164 |  |  |  | 0.1 | 0.84034
13 | 10.8 | 0.95359 | 10.9 | 0.93142 |  |  |  | -0.1 | -0.92593
14 | 11.4 | 1.0093 | 11.2 | 0.95966 |  |  |  | 0.2 | 1.7544
15 | 11 | 0.97214 | 11 | 0.94083 |  |  |  | 0 | 0
16 | 11.8 | 1.0465 | 11.6 | 0.99743 |  |  |  | 0.2 | 1.6949
17 | 11.7 | 1.0372 | 11.9 | 1.0258 |  |  |  | -0.2 | -1.7094
18 | 11.4 | 1.0093 | 11.5 | 0.98798 |  |  |  | -0.1 | -0.87719
19 | 12.7 | 1.1305 | 12.6 | 1.0924 |  |  |  | 0.1 | 0.7874
20 | 12.6 | 1.1211 | 12.5 | 1.0829 |  |  |  | 0.1 | 0.79365
New Geometric Mean |  |  | 11.6272
Current Refrence Range Validated? |  |  |  | Yes
If no, complete page 2
Comments: Acceptable.
Slope: | 1.034
R: | 0.97274
Intercept: | -0.37523
Supervisor Review EL 1/5/07
Page 2: Updated Reference Ranges
|  | Date |  | Analyzer |  | New Lot#
|  | Tech |  | Serial Number |  | Expiration
| Prothromin Time |  |  |  | INR
| Old Lower Limit |  | Old Upper Limit |  | Old Lower Limit |  | Old Upper Limit
| New Lower Limit |  | New Upper Limit |  | New Lower Limit |  | New Upper Limit
New Adverse Event Limits
|  | PT | INR |  | Results greater than the values listed are Adverse Events as graded
| Grade 1 | 0 | 0
| Grade 2 | 0 | 0
| Grade 3 | 0 | 0
| Grade 4 | 0 | 0
Supervisor Review